Maxim WELLNESS APP USER GUIDE & WellNESS LIBRARY INTEGRATION GUIDE

Version 4.4.5

January 14, 2022

Maxim corner motif.emf

Contents

[1 Introduction 4](#_Toc80370134)

[2 Wellness Suite Application 5](#_Toc80370135)

[2.1 System Architecture 5](#_Toc80370136)

[2.2 Application Details 5](#_Toc80370137)

[2.2.1 Permissions 6](#_Toc80370138)

[2.2.2 Discover BLE Devices 7](#_Toc80370139)

[2.2.3 Landing Page 8](#_Toc80370140)

[2.2.4 User Profile 8](#_Toc80370141)

[2.2.5 Optical HRM 11](#_Toc80370142)

[2.2.6 Pulse Oximetry 13](#_Toc80370143)

[2.2.7 Heart Rate Variability 14](#_Toc80370144)

[2.2.8 Respiratory Rate 16](#_Toc80370145)

[2.2.9 Stress 17](#_Toc80370146)

[2.2.10 Archive 18](#_Toc80370147)

[2.2.11 Sleep Quality 20](#_Toc80370148)

[2.2.12 Sports Coaching 22](#_Toc80370149)

[2.2.13 Flash Log Parser 26](#_Toc80370150)

[2.2.14 HR Alignment 27](#_Toc80370151)

[3 Wellness Library Configuration Definitions 29](#_Toc80370152)

[3.1 HRV Algorithm Configuration 29](#_Toc80370153)

[3.2 Respiration Rate Manager Algorithm Configuration 30](#_Toc80370154)

[3.3 Sleep Quality Assessment Algorithm Configuration 30](#_Toc80370155)

[3.4 Stress Monitoring Algorithm Configuration 31](#_Toc80370156)

[4 Wellness Library I/O Interface Attribute Definitions 32](#_Toc80370157)

[4.1 HRV Algorithm’s Output and Status Definitions 33](#_Toc80370158)

[4.2 Respiration Rate Algorithm’s Output and Status Definitions 34](#_Toc80370159)

[4.3 Sleep Quality Assessment Algorithm’s Output and Status Definitions 34](#_Toc80370160)

[4.4 Stress Monitoring Algorithm’s Output and Status Definitions 36](#_Toc80370161)

[5 Wellness Library API Function Definitions 37](#_Toc80370162)

[5.1 Algorithm Related Process 37](#_Toc80370163)

[5.2 Authentication Related Process 38](#_Toc80370164)

TABLEs

Table 1: Stress Score 18

Table 2: Stress Score Provision 36

Figures

Figure 1: System Architecture 5

Figure 2: Android App Flow Diagram 6

Figure 3:Permissions 7

Figure 4: Available Devices 7

Figure 5: Landing Page 8

Figure 6: User Profile Page 9

Figure 7: Optical HRM 11

Figure 8: Optical HRM Menu 12

Figure 9: Pulse Oximetry 13

Figure 10: Pulse Oximetry – User Info 14

Figure 11: Heart Rate Variability 16

Figure 12: Respiratory Rate 17

Figure 13: Stress 18

Figure 14: Archive 19

Figure 15: Archive – Preview Screen 20

Figure 16: Sleep Quality Assessment 21

Figure 17: Sleep Quality Assessment-Detailed Page 21

Figure 18: Sports Coaching Landing Page 22

Figure 19: Sports Coaching Training page 23

Figure 20: Readiness 23

Figure 20: Resting VO2Max 24

Figure 21: EPOC Recovery 24

Figure 22: Recovery Time 25

Figure 23: Fitness Age 25

Figure 24: Sports Coaching User History 26

Figure 25:Flash Log Parser 26

Figure 26: HR Alignment 27

# Introduction

The purpose of the Wellness Suite application is to provide a single Android GUI to evaluate all algorithms in Maxim portfolio. Algorithms consuming outputs of the ones running on the MRD103/4/5 modules will be executed on the Android platform. From the user perspective, the Wellness Suite is the interface to all Maxim algorithms.

This user guide covers these algorithms: WHRM, SpO2, BPT, HRV, Respiration Rate, Stress, Sleep and Sports Coaching. The first three of them runs by the MRD103/4/5, the rest runs by the mobile application. In addition to these algorithms, the application provides auxiliary functions such as Flash Log Parser and HR Alignment.

**NOTE: The minimum supported Android version is 5, the targeted Android version is Android 11.**

# Wellness Suite Application

## System Architecture

The mobile application will communicate with the MAXREFDES103/4/5 via BLE and feed the bio-sensing algorithms included in the application. The complete architecture can be viewed in Figure 1.




Figure : System Architecture

## Application Details


Wellness Suite application runs on Android operating systems. The minimum supported android version for this application is Android 5.0. Besides, the mobile phone running the application should have at least BLE version 4.2. The app flow diagram can be seen in Figure 2.  
 

Figure : Android App Flow Diagram

### Permissions

After opening the application for the first time, the application asks the user to grant the necessary permissions to run as it is supposed to. The application requires the below permissions:

* Location Access
* Storage Write
* Storage Read
* Enabling Bluetooth

The location access permission is needed for the application to discover the nearby Bluetooth Devices. The Storage Write/Read permissions are needed to log the sensor data and read them back. The user should also enable the Bluetooth to discover devices and communicate with them.




  

Figure :Permissions

### Discover BLE Devices

After granting permissions, the available Maxim devices are listed automatically. The user can start the scanning and stop it by pressing the button located in upper-right position.

Every item in the list shows the device name, device MAC address and its signal strength.

**NOTE: For some mobile phones, the GPS should be open to discover nearby BLE devices. If no device is listed in the app, please open the GPS to see the available devices.**




Figure : Available Devices

When the user selects the appropriate device by pressing on it, the application shows the Landing Page (Main Activity) to the user.

User should be connected to the device to be able to start any measurement. If device disconnect any time, user will be warned by a dialog like in Figure 5.




Figure 5: Connection Lost Warning Dialog

### Landing Page

When the user arrives to this page, the application is already connected to the selected device and starts communicating. It receives the device information including hub version. On the upper-left corner, Bluetooth icon can be seen. It indicates the connection status of the selected device. On the lower-left corner, the app version is shown. On the lower-right corner, the device name and the hub version are shown.



 

Figure : Landing Page

This page contains buttons directing user to the related pages. For example, when the User Profile button is pressed, User Profile page opens.

### User Profile

The algorithms that are developed for both the watch itself and Android require a few user information to give more concise result. These user data can be cited as birth year, weight, height, initial heart rate, sleep resting heart rate and the gender. In addition to this information, the application also needs the username to maintain different user profiles and allow users to select their own personal information dataset.



 

Figure : User Profile Page

In the first page of the User Profile, all the available users are listed. The selected user is differentiated from others by its green header and the check mark. The selected user profile is used for both the device algorithms and the android algorithms. If none of the available users is selected, the default user profile is going to be used for them.

With the floating action button located at bottom right of the page, user can create a new user profile.

**NOTE: This collected data is used to improve the accuracy of Maxim algorithms that run on this device. It is not shared with any 3rd parties.**

### Blood Pressure Trending

Blood pressure (BP) is a vital sign that provides critical information about cardiac output, cardiac pressure, and hemodynamic stability. BP correlates well with the risk of advanced cardiovascular disease (ACVD). Also, hypertension, a medical condition in which systolic BP is above 140mmHg and/or diastolic BP is above 90mmHg persistently, is a risk factor for coronary heart disease and the single most important risk factor for stroke. Systolic BP is the maximum pressure in the arteries when the heart contracts. Diastolic BP is the minimum pressure in the arteries between the heart's contractions.

Therefore, continuously monitoring BP is valuable. However, the conventional noninvasive method—a cuff sphygmomanometer—usually allows only one reading every few minutes. Moreover, the inflation process disturbs the user and makes BP measurement inconvenient during sleep or exercise.

Maxim Blood Pressure Trending (MaximBPT) is a BP measurement system using a photoplethysmography (PPG) signal obtained from the fingertip. The system requires user-specific calibrations, and successive measurements are relative to this calibration point. The goal of the system is to make BP measurements readily available to the user through widely available consumer devices like smartphones or wellness watches. The system is most useful and accurate when measurements are taken in a resting state, as in traditional cuff-based BP systems.

In the landing page of BPT, users can connect with BPT reference device which is Qardio.




 

Figure 8: BPT Landing Page

If BPT reference device connection made, calibration reference values are taken automatically. If BPT reference device is not connected, users should enter the reference values manually. For the sake of measurement, users should rest 5 minutes before calibration.

Graphical user interface, application

Description automatically generatedGraphical user interface, text, application



Description automatically generatedGraphical user interface, application

Description automatically generatedGraphical user interface, application

Description automatically generated

Figure 9: BPT Calibration

Also in the measuring part, if BPT reference device connected, users can compare the results. For the sake of measurement, BPT algorithm run before reference device.

Graphical user interface, application

Description automatically generatedA picture containing chart


Description automatically generated

Figure 10: BPT Measurement

Users can view the history through landing page.




Figure 11: BPT History Page

**NOTE: Only MRD104 has support BPT feature.**

### Optical HRM

Heart rate is the frequency of heart contractions that is measured in beats per minute (BPM). Maxim Wearable Heart Rate Monitoring (WHRM) algorithm utilizes synchronized 3-axis accelerometer and PPG data to provide several end-user friendly fitness tracking outputs like heart rate and activity related statistics e.g. step count, burned calories and activity class of the user. The PPG (photoplethysmography) signal can be gathered from contractions of ventricles in finger, wrist, ear, toe, chest etc.

Maxim WHRM solution includes HR and activity tracking algorithms, especially for the heart rate detection, cardiac beat detection, step measurement and activity classification. The envisioned product targets health and wellness applications for continuous 24/7 monitoring with optimized power management and the provision of accurate HR and IBI detection, step detection and sample by sample activity classification.

As seen in the Figure 12, this page is responsible to show the WHRM Outputs such as Heart Rate (HR) and activity related outputs. The application also shows the PPG data on the chart.







Figure : Optical HRM

From the upper right corner, the user can enable/disable the features such as “Log to File”, “Log to Flash”, “Enable SCD”, “Enable MFIO”, “Enable SCD SM”, “Low Power Mode”. Also, the user can open the settings.

* Log to File: If it is enabled, the received sensor data is recorded and saved to a file under “MaximSensorsApp/RAW” folder. The app also logs a 1 Hz HR data and saved to a file under “MaximSensorsApp/1Hz” folder.
* Log to Flash: If it is checked, the device stores the data in its flash, and it does not send the data to the app.
* Enable SCD: If it is checked, the algorithm detects if the watch is worn.
* Enable MFIO: If it is checked, the app enables reporting of events from SensorHub via interrupt.
* Enable SCD SM: If it is checked, Skin Contact Detection State Machine is enabled.
* Low Power Mode: If it is checked, the data stream rate becomes 1Hz instead of 25Hz. Internally, we are generating 25Hz data from 1Hz data to support logging and algorithm functionality.

From the settings dialog (Figure 13), the user can change the following:

* Sample mode cycle time
* HR Confidence Level
* HR Expiration Threshold

The algorithm has 2 modes: Continuous mode and Sampled mode. In the continuous mode, the data streaming does not stop until the user stops it. In the sampled mode, the algorithm stops when the calculation finished, and restarts again after certain time. This time is determined from the Settings Dialog as seen in Figure 13. According to HR Confidence Level and HR Expiration Threshold settings, If the low confidence condition is met, the app shows a warning to the user as seen in Figure 13.




  

Figure : Optical HRM Menu

The user starts and stops monitoring by pressing the play and stop button, respectively. Also, he/she can add an annotation while measuring. It will be saved to “MaximSensorsApp/ANNOTATIONS” folder.

While receiving data from the device, if packet loss occurs, the app records an error log file under “MaximSensorsApp/ERROR” together with a warning message.

This page also allows user to connect reference HR device to compare its results with Maxim’s results. The received data is recorded and saved to a file under “MaximSensorsApp/HR\_REF” folder. After stopping the measurement, the app tries to align the Maxim HR results with Reference HR data. If the alignment is done successfully, the app saved the aligned HR Results to a file under “MaximSensorsApp/ALIGNED” folder.

**Note:** The Annotation feature and Error log feature are available for all the measurements.

### Pulse Oximetry

Oxygen plays vital role for every cell in human body to produce energy and all organs require oxygen for metabolism. Pulse oximetry is an essential technique used for many decades to monitor the amount of oxygen carried in blood. Maxim has developed high-sensitive finger-based optical sensors with optimized algorithm to make the pulse oximeters portable and easily applicable.

Pulse oximeter is a lightweight, small and noninvasive optical devices to monitor the amount of oxygen carried in the blood. Pulse oximeter devices measure oxygen saturation by illuminating the skin with red and near-infrared (IR) lights and measure changes in light absorption of oxygenated (oxy-hemoglobin, HbO2) and deoxygenated blood (deoxy-hemoglobin, RHb)) using two light wavelengths.

As seen in the Figure 14, this page is responsible to show the SpO2 value together with Heart Rate. The application also shows the PPG data on the chart.

The algorithm has 2 modes: Continuous mode and One-shot mode. In the continuous mode, the data streaming doesn’t stop until the user stops it. In the one-shot mode, the algorithm stops when the calculation finished.

The user starts and stops monitoring by pressing the play and stop button, respectively.



 

Figure : Pulse Oximetry

The SpO2 calculation requires “no movements and talk”. When the user measures his/her SpO2 rate, he/she needs to stay still and should not talk. The application gives visual indication for the parameters that affect the measurement such as Signal Quality, Motion, and Orientation. If these indicators are red, it means there is a problem in the corresponding parameter.

This page also allows user to connect reference SpO2 device to compare its results with Maxim’s results. The received data is recorded and saved to a file under “MaximSensorsApp/SPO2\_REF” folder.

The user can log user information such as ID, Name, Surname, Birth Year, Weight, Height, Hemoglobin if “Log user Info” option is selected. When “Log User Info” is enabled, text boxes, that can be seen in Figure 15, must be entered, and log file has ID info as appended to the end of filename. Other data will be stored in the folder named USER\_INFO.



 

Figure 15: Pulse Oximetry – User Info

**NOTE: The user should adjust the wristwatch on the wrist such that Red PPG peaks and troughs are clearly visible on the chart. After every valid SpO2 value, the color text that represents the SpO2 value is changed to black, and then it fades out until a new value is received. If a new SpO2 measurement is not calculated in 10 seconds after the last measurement, the value becomes obsolete and app shows “No Report” instead of old value.**

### Heart Rate Variability

Heart Rate Variability (HRV) provides a measure of the time difference between successive heartbeats; the variation is controlled by the autonomic nervous system (ANS). This is the system that regulates our heartbeat, blood pressure, and breathing. Our ANS consists of two main components:

* Our “fight-or-flight mechanism” (the sympathetic nervous system) is our response to external stress factors. This is what makes us energized and ready to face challenges.
* Our relaxation response (the parasympathetic nervous system) is the opposite. This is what turns us toward a rest and recovery mode.

HRV tends to be low when we are in a fight-or-flight mode and high when we are more relaxed. High HRV is generally associated with greater cardiovascular fitness. The electrocardiogram test is the gold standard for assessing HRV. However, wearable devices with sophisticated algorithms are now making this data more accessible to more people. Since HRV provides insights into our ANS, which regulates our heart rate (HR) and blood pressure, it can be considered as the most important general wellness parameter.

HRV gives us some clues about potential cardiac conditions. Over time, higher HRV can indicate increased resilience, while lower HRV can point to chronic stress. HRV can be increased by changes in behavior: quitting smoking, losing weight if needed, exercising, and managing stress, for example. HRV analysis techniques that fall into three categories provide useful health insights:

* **Time-domain** analysis based on standard deviation and root mean square (RMS) of normal-to-normal intervals provides powerful differentiators of stressed versus relaxed conditions. Increased HRV is an indicator of fitness.
* **Frequency domain** analysis shows the ratio of parasympathetic and sympathetic activity.
  + High frequency pertains to the parasympathetic system and the vagus nerve, which controls the parasympathetic nervous system
  + Low frequency pertains to sympathetic activity
  + Very low frequency pertains to the sympathetic nervous system, chemoreceptors, thermoreceptors, and the renin-angiotensin system (i.e., hormones)
* **Non-linear** analysis can point to underlying cardiac conditions and includes:
  + Detrended fluctuation analysis (DFA), which looks for self-similar patterns by analyzing the power spectral density (PSD). Peaks in PSD indicate repetitive patterns.
  + Entropy analysis, a measure of randomness over time. Decreased HRV and increased randomness of HR are independently associated with high-risk conditions.
  + Poincare plot analysis utilizes scatter plots of consecutive pulse interval points. Consecutive pulses that vary by large amounts will have larger scattering around the diagonal. Low HRV will shrink and cluster around the diagonal. Unbalanced HR behaviors such as fast acceleration and slow deceleration will generate asymmetric plots. Large off-diagonals show skipped heartbeats, which usually indicate an arrhythmia problem.

As seen in the Figure 16, this page is responsible to show the HRV algorithm results together with IBI values. The application receives the data stream and passes them to the algorithm library. Then, it displays the algorithm results.

The meanings of the HRV metrics are listed below:

* **Time Domain Metrics**
  + **AVNN:** Average of normal to normal (NN) intervals in ms
  + **SDNN:** Standard deviation of NN intervals in ms
  + **RMSSD:** RMS value of successive differences in ms
  + **PNN50:** Percentage of successive differences greater than 50 ms
* **Frequency Domain Metrics**
  + **ULF:** Power in Ultra Low Frequency band in ms2
  + **VLF:** Power in Very Low Frequency band in ms2
  + **LF:** Power in Low Frequency band in ms2
  + **HF:** Power in High Frequency band in ms2
  + **LF/HF:** LF/HF ratio
  + **TOTPWR:** Total power in ms2





Figure : Heart Rate Variability

### Respiratory Rate

Respiration rate is a vital sign indicating health and wellness and can be an input for several applications such as sleep quality assessment, stress level estimation and energy exposure estimation etc. The available methods to measure respiration rate base on capnography, ECG and PPG. Although capnography has the most precise results, it needs oral/nasal cannula which makes impossible to use it under movement. ECG based solutions need multiple probs at different locations on human body, therefore they are also not convenient for a watch-based design. Maxim’s solution is based on PPG and compatible to use on wrist, finger and forehead that make it best solution for wearable devices.

Using PPG, we can measure respiration induced features. One of these features is Respiratory Sinus Arrhythmia (RSA) which causes heart rate change during respiratory events. Another feature we can measure is that the effect on the venous volume of intrathoracic pressure changes during respiratory cycle. After processing these features, our solution can report very accurate respiration rate estimates.

As seen in the Figure 17, this page is responsible to show the Respiration Rate Algorithm result. The application also shows the Respiration Rate on the chart. The application receives the data stream and passes them to the algorithm library. Then, it displays the algorithm results.

This page also allows user to connect reference RR device to compare its results with Maxim’s results. The received data is recorded and saved to a file under “MaximSensorsApp/RR \_REF” folder.





Figure : Respiratory Rate

### Stress

Daily life stress is one of the important issues of modern life. Mainly, two kinds of stress phenomenon exist as acute and chronic stress. American Psychological Association noted that the causes of acute stress are pressure from recent past and near future. In example, exercise challenges or any kind of sudden anxiety can induce acute stress. On the other hand, chronic stress results from the long-term pressures like socioeconomic conditions, ongoing problems in relations etc. The scope of this study is focused on the detection of acute stress.

Stress releases a hormone called adrenaline that quickens breath increases HR and Blood Pressure, known as "fight or flight response". Deep breathing exercise therefore can be suggested to shift body and mind into more relaxed state.

As there is no single gold standard for measuring acute mental stress, several measurement methods were employed in related works like: psychological tests, blood tests or salivary cortisol level tests. In literature, most of the wearable stress monitoring studies use galvanic skin response, muscle tension, blood pressure and skin temperature devices as well as HR monitoring devices. In this study, Maxim’s stress estimation methods use accelerometer and PPG data (HR, HRV) solely.

Quick Stress Assessment for acute stress is in the scope of this solution where user provides a considerably small amount of data (using wearable device) in a resting state and receives his stress level information.

As seen in the Figure 18, this page is responsible to show Stress Algorithm result together with the HR value. The application receives the data stream and passes them to the algorithm library. Then, it displays the algorithm results.




Figure : Stress

The stress score is in the range of 1 to 18.

|  |  |  |  |
| --- | --- | --- | --- |
| **Score** | **Meaning** | **Color of the Emoji** | **Status of the Emoji** |
| 1-2 | Overwhelmed | RED | ☹ |
| 3-5 | Frustrated | ORANGE | ☹ |
| 6-7 | Manageable Stress | ORANGE | 😐 |
| 8-9 | Feeling Okay | GREEN | 😐 |
| 10-11 | Doing Great | GREEN | 😊 |
| 12-14 | Very Relaxed | ORANGE | 😊 |
| 15-18 | Uncomfortably Relaxed | RED | 😊 |

Table : Stress Score

### Archive

The archive page lists all the log files. It reads the file names and parses them to show to the user in a more user-friendly way.

If no user is selected in the Main Landing Page, the app does not allow user to continue to Archive page. If the user already selected the profile, the app navigates the user to the Archive page.





Figure : Archive

When the user presses to list item, the application reads the recorded sensor data and passes them to the algorithm library. Every list item contains the below buttons:

: It deletes the related file.  
: It allows user to share the file.  
: It runs the sleep algorithm on the related file. After calculation, the app warns the user whether the algorithm finds sleep data in the log file or not. If it finds, the app saves the sleep algorithm result under “MaximSleepQA” folder.

: It allows user to add blood gas measurement.

As seen in Figure 20, the user can select the graph from the Spinner to visualize it. The spinner contains: HR, SpO2, IBI, Steps, Motion, Resp. Rate, HRV metrics and Stress Score. If the selected graph is the result of an algorithm running in the app, the Export CSV button is going to be available for the user to export the algorithm result as csv.




  

Figure : Archive – Preview Screen

### Sleep Quality

Sleep is a crucial element for human’s mental and physical. It is a highly active state of the body where memory formation and restoration, removal of the unimportant connections and damages during wake state such as oxidative stress, death of neurons in the hippocampus etc. are held. As a result, a healthy sleep is one of the key elements to able to use the cognitive skills. It is stated that not only lack of sleep cause severe losses in cognitive abilities, but also cause cardiovascular diseases.

Sleep-Wake cycle is dictated by the hippocampus through the circadian rhythm. This process stimulates the sleep-promoting and arousal centers in the brain through neurochemical processes. Sleep is characterized by the cyclic occurrences of two distinct states known as rapid eye movement (REM) and non-rapid eye movement (NREM). Furthermore, the NREM consist of slow wave sleep so called Deep Sleep, and lighter sleep stages so called Light Sleep. During the REM dreams occur and activity of the brain excites like the wake state as of other biological traits. As the name suggests, during Deep Sleep brain reflects slower.

The gold standard for the tracking of known as polysomnography (PSG). PSG generally comprises a human annotator, classifying the non-overlapping epochs of 30s to a sleep stage via analyzing electroencephalogram (EEG), electrocardiogram (ECG), electromyogram, electrooculogram, oxygen saturation, respiratory effort, and movement signals. However, not only PSG requires expensive laboratory facilities but also disrupts the sleep of the subjects due to bulky devices. Maxim Integrated has developed an affordable and non-disruptive sleep tracking solution for wrist worn devices.

The Sleep Quality Assessment application is integrated into Wellness Suite App v3.0.0. The app looks under the “MaximSleepQA” folder to visualize the results.

If no user is selected in the Main Landing Page, the app does not allow user to continue to Sleep Quality. If the user already selected the profile, the app navigates the user to the Sleep Quality.

In the main page of the Sleep Fragment, the bar chart lists the sleep time date by date. Below the bar chart, there are pie charts showing the stages of the sleep data for the day in which the measurement is taken.




Figure : Sleep Quality Assessment

The “Details” button takes the user to a more detailed page about the corresponding sleep data. At that page, the user can see his/her sleep score, sleep stage durations, movements, Heart Rate and IBI values.







Figure : Sleep Quality Assessment-Detailed Page

### Sports Coaching

To achieve a healthy life, sport needs to be a part of the individuals lives. People need to plan and keep track of their sports performance and frequency to adjust their lifestyle, so that they can keep their healthy state. With the widespread of wearable devices, people started to use them for keeping track of their vital sign (heart rate, SpO2, etc.) and sports activities (walking, running, biking, etc.). As a next step, people are requesting to keep track of their performance and effort using their wearable devices. Maxim developed bunch of the algorithms to assess users’ performance, effort, readiness during sport activities and bundled them in the Maxim Sports Coaching Algorithms product. Users can enhance their performance and decrease sports related injuries using Maxim Sports Coaching Algorithms.

#### Landing Page

Maxim sports coaching algorithms require the users’ personal information such as birth year, gender, weight, and height for the measurements. If no user is selected in the Main Landing Page, the app does not allow user to continue to Sports Coaching session.

If the user already selected the profile, the app navigates the user to the Sports Coaching Landing page. From this page user can do the followings:

* Training: User can visit all the sports coaching measurements pages.
* History: User can see his/her measurement records.
* Profile: User can change his/her profile.




Figure : Sports Coaching Landing Page

#### Training

This page lists all the measurements that the Maxim Sports Coaching Algorithm supports. As seen in Figure 24, the user can measure his/her Readiness, Resting VO2Max, EPOC Recovery, Recovery Time, and Fitness Age.




Figure : Sports Coaching Training page

#### Readiness

This function is used to estimate how much ready is the user for the next training. It needs 3-6 minutes of data collected from the user in resting state. User can start measuring by pressing the play button. While measuring, the completion percentage is reported in the circular progress bar, and the heart rate is displayed in real-time. When the measurement finishes, the user can see his readiness score together with HR statistics as seen in Figure 25.




Figure : Readiness

#### Resting VO2Max

This function is used to calculate VO2\_MAX of the user. This function needs 3-6 minutes of data collected from the user while they are sitting in a relax state. Relax state is where user is fully recovered from last training and not effected by any kind of food, drug, or alcohol. User can start measuring by pressing the play button. While measuring, the completion percentage is reported in the circular progress bar, and the heart rate is displayed in real-time. When the measurement finishes, the user can see his VO2Max score together with HR statistics as seen in Figure 26.




Figure : Resting VO2Max

#### EPOC Recovery

EPOC is an abbreviation for Excess Post-Exercise Oxygen Consumption. To estimate the user’s EPOC value, user needs to start the algorithm after training and collect data for 5-6 minutes in resting state. User needs to select the appropriate parameters before starting the measurement. This functionality is dependent on the user information such as exercise duration in minutes, exercise intensity and the time passed since the end of exercise.



 

Figure : EPOC Recovery

After selecting these parameters, user can start measuring by pressing the play button. While measuring, the completion percentage is reported in the circular progress bar, and the heart rate is displayed in real-time. When the measurement finishes, the user can see his EPOC score together with HR statistics as seen in Figure 27.

**Note 1:** Mean HR is not reported for EPOC. It is a known issue and will be fixed in the next release.

**Note 2:** To calculate the EPOC Recovery, there should be at least 1 VO2Max calculation in the history.

#### Recovery Time

Recovery Time is the amount of time until user reaches to their homeostasis. Algorithm needs to gather 2-3 minutes of data from the user to estimate his/her recovery time. User can start measuring by pressing the play button. While measuring, the completion percentage is reported in the circular progress bar, and the heart rate is displayed in real-time. When the measurement finishes, the user can see his recovery time together with HR statistics as seen in Figure 28.

**Note 1:** Mean HR is not reported for Recovery Time. It is a known issue and will be fixed in the next release.

**Note 2:** To calculate the Recovery Time, there should be at least 1 EPOC calculation in the history.




Figure : Recovery Time

#### Fitness Age

This page calculates the user’s fitness age from the history data and compare it with his/her actual age as seen in Figure 29.




Figure : Fitness Age

#### History

The History page shows the list of all the measurements that the user takes. This page is searching for the files under “MaximSensorsApp/SPORTS\_COACHING/<username>/” folder. Every item displays the date the measurement is taken, session name and score as seen in Figure 30.




Figure : Sports Coaching User History

### Flash Log Parser

Logging the sensor data is one of the crucial functions of the wearable devices. MRD103 devices are not only able to send the sensor data to the other devices but also save these data in its flash. The user can obtain these data by turning the mode of the device into storage mode. The extension for these files is “maximlog”.

To visualize this log file in Wellness Suite app, it needs to be converted to a csv file with a predefined format. Wellness Suite app provides Flash Log Parser functionality to the user.




Figure :Flash Log Parser

As seen in Figure 31, the user needs to select the log file by pressing the “IMPORT FILE” button. He/she can see the log file location in the screen. To parse the log file, the user should press the “PARSE” button. After parsing is finished, the location to the csv file is displayed in the screen. If any problems occurred during parsing, the output file is not created, and the error message is displayed to the user.

**Note: MRD104 devices do not support this feature.**

### HR Alignment

Wellness Suite app supports automatic Heart Rate Alignment. If the user wants to compare the Maxim heart rate results with a reference device’s results, he/she can use this functionality of the Wellness app. There are 3 restrictions to accomplish this:

1. The MRD103 and the reference device should have been used at the same time.
2. Maxim raw data should be used for the alignment.
3. Reference HR log file should contain 2 columns: UNIX timestamp (ms) and Heart Rate (bpm). See Figure 32.

As seen in Figure 32, the user needs to select the Maxim Raw Data File by pressing the corresponding Attachment icon. After selecting the file, the file location is displayed in the screen. Similarly, the user should also select the reference HR file. After selecting both files, the user can align them automatically by pressing the “ALIGN” button. The Wellness Suite app displays the status of the alignment process.



 

Figure : HR Alignment

# Wellness Library Configuration Definitions

Wellness Library offers single configuration structure which covers configuration parameters for all algorithms included in the application. The config structure is defined in “Wrapper.h file”:

**struct** mxm\_algosuite\_init\_data {

**unsigned** **char** enabledAlgorithms;

MxmHrvConfig hrvConfig;

mxm\_respiration\_rate\_manager\_init\_str respConfig;

mxm\_stress\_monitoring\_config stressConfig;

mxm\_sleep\_manager\_config sleepConfig;

};

*enabledAlgorithms* field is the 8bit long field where Algorithms to be initialized at configuration is selected. Selection is done via configuration bits dedicated to each underlying algorithm:

**#define** MXM\_ALGOSUITE\_ENABLE\_HRV ( 1 << 0)

**#define** MXM\_ALGOSUITE\_ENABLE\_RESP ( 1 << 1)

**#define** MXM\_ALGOSUITE\_ENABLE\_SLEEP ( 1 << 2)

**#define** MXM\_ALGOSUITE\_ENABLE\_STRESS ( 1 << 3)

**#define** MXM\_ALGOSUITE\_ENABLE\_SPORTS ( 1 << 4)

Selected algorithms are initialized within the initialization phase of library and driven by inputs during running phase.

Each underlying library has its dedicated configuration field within the main configuration structure as define below.

## HRV Algorithm Configuration

Heart Rate Variation Calculation Algorithm’s configuration field is described in “mxm\_hrv\_public.h“ API file distributed with Maxim Wellness Library. It is configuration structure is defined as:

**struct** \_MxmHrvConfig {

**float** samplingPeriod;

uint16\_t windowSizeInSec;

uint16\_t windowShiftSizeInSec;

}

*SamplingPeriod*: Input sample frequency of the HRV algorithm. It is recommended to set as 40.0 ms when other algorithms within the Library is enabled to run.

*windowSizeInSec*: Size of the sample buffer processed by HRV algorithm. Determines number of input samples to make an estimation. In sample’s unit it is windowSizeInSec/ samplingPeriod.

windowShiftSizeInSec: Determines resolution of HRV outputs. Smaller in the shift means increase in resolution. In sample’s unit it is windowShiftSizeInSec / samplingPeriod.

## Respiration Rate Manager Algorithm Configuration

Respiration Rate Monitoring Algorithm’s configuration field is described in “mxm\_respiration\_rate\_manager.h “ API file distributed with Maxim Wellness Library. It is configuration structure is defined as:

**struct** \_mxm\_respiration\_rate\_manager\_init\_str {

mxm\_respiration\_rate\_manager\_ppg\_source\_options signal\_source\_option;

mxm\_respiration\_rate\_manager\_led\_codes led\_code;

mxm\_respiration\_rate\_manager\_sampling\_rate\_option sampling\_rate;

uint32\_t motion\_magnitude\_limit;

};

*signal\_source\_option*: setting for the body part where PPG signal is obtained

*led\_code*: selection of LED source PPG data is excited from. Used fr identifying correct array indices for multiple LED channels

*sampling rate*: input sample rate selection. 25Hz and 100Hz options are accepted. If other algorithms are enabled to run in library, it is recommended to keep in 25Hz to get correct results or check for the sampling rate of other enabled algorithms within library.

*motion\_magnitude\_limit*: The motion magnitude limit in milig.

## Sleep Quality Assessment Algorithm Configuration

Sleep Quality Assessment Algorithm’s configuration field is described “mxm\_sleep\_manager.h” file distributed with Maxim Wellness Library. It is configuration structure is defined as:

**struct** \_mxm\_sleep\_manager\_config

{

mxm\_sleep\_manager\_minimum\_detectable\_sleep\_duration mxm\_sleep\_detection\_duration;

mxm\_sleep\_user\_info user\_info;

**bool** is\_resting\_hr\_available;

**bool** is\_confidence\_level\_available\_hr;

**bool** is\_confidence\_level\_available\_ibi;

**bool** is\_activity\_available;

} mxm\_sleep\_manager\_config;

*mxm\_sleep\_detection\_duration*: sets the minimum time sleep algorithm will start to produce accurate classifier results.

*user\_info*: age, weight, gender and sleep resting heart rate of user

*is\_resting\_hr\_available*: set weather value of sleep resting heart rate of user is input during configuration

*is\_confidence\_level\_available\_hr*: Availability of the confidence level for the Heart Rate Measurement

*is\_confidence\_level\_available\_ibi* : Availability of the confidence level for the Interbeat Interval Measurement

*is\_activity\_available*: Availability of the activity class identifier

## Stress Monitoring Algorithm Configuration

Stress Monitoring Algorithm’s configuration field is described “mxm\_stress\_monitoring.h” file distributed with Maxim Wellness Library. It is configuration structure is defined as:

**struct** \_mxm\_stress\_monitoring\_config {

uint8\_t dummy\_config\_for\_compilation;

}

Algorithm’s configuration uses a dummy configuration object to obey Maxim Wellness Library algorithm’s notation.

# Wellness Library I/O Interface Attribute Definitions

Maxim Wellness Library input format is defined in “AlgoWrapper.h” file distributed with Maxim Wellness Library:

**struct** mxm\_algosuite\_input\_data {

uint32\_t inp\_sample\_count ;

uint32\_t grn\_count ;

uint32\_t grn2Cnt ;

uint32\_t irCnt ;

uint32\_t redCnt ;

uint32\_t accelx ;

uint32\_t accely ;

uint32\_t accelz ;

uint32\_t whrm\_suite\_curr\_opmode ;

uint32\_t hearth\_rate\_estim ;

uint32\_t hr\_confidence ;

uint32\_t rr\_interbeat\_interval ;

uint32\_t rr\_confidence ;

uint32\_t activity\_class ;

uint32\_t r\_spo2 ;

uint32\_t spo2\_confidence ;

uint32\_t spo2\_estim ;

uint32\_t spo2\_calc\_percentage ;

uint32\_t spo2\_low\_sign\_quality\_flag ;

uint32\_t spo2\_motion\_flag ;

uint32\_t spo2\_low\_pi\_flag ;

uint32\_t spo2\_unreliable\_r\_flag ;

uint32\_t spo2\_state ;

uint32\_t skin\_contact\_state ;

uint32\_t walk\_steps ;

uint32\_t run\_steps ;

uint32\_t kcal ;

uint32\_t cadence ;

uint32\_t timestampUpper32bit ;

uint32\_t timestampLower32bit ;

};

Input of the Library is based on the extended report output of the WHRM suite running within SensorHub processor, i.e. ME11. **For the definition of field within input structure, please refer to API document of Max32664C.**

**IMPORTANT: Input sampling frequency must be a multiple of 25 samples per second. 25Sps is the preferred rate.**

Maxim Wellness Library output format is defined in “AlgoWrapper.h” file distributed with Maxim Wellness Library:

**Struct** mxm\_algosuite\_output\_data {

MxmHrvOutData hrv\_out\_sample;

mxm\_sleep\_manager\_output\_dataframe sleep\_out\_Sample;

mxm\_respiration\_rate\_manager\_out\_data\_str resp\_out\_sample;

mxm\_mxm\_stress\_monitoring\_run\_output stress\_out\_sample;

};

Where each algorithm’s output has reserved fields.

Maxim Wellness Library return codes are defined in “AlgoWrapper.h” file distributed with Maxim Wellness Library:

**Struct** mxm\_algosuite\_return\_code {

mxm\_algosuite\_hrv\_retcode hrv\_status;

mxm\_algosuite\_resp\_retcode resp\_status;

mxm\_algosuite\_stress\_retcode stress\_status;

mxm\_algosuite\_sleep\_retcode sleep\_status;

}

Where each algorithm’s status/return reporting is done in its reserved field.

## HRV Algorithm’s Output and Status Definitions

HRV Algorithm’s output field is defined in “mxm\_hrv\_public.h” file distributed with Maxim Wellness Library:

**struct** \_MxmHrvOutData {

TimeDomainHrvMetrics timeDomainMetrics;

FreqDomainHrvMetrics freqDomainMetrics;

**int** percentCompleted;

**bool** isHrvCalculated;

}

*timeDomainMetrics*: this is time domain related output field containing following attributes:

* + **AVNN:** Average of normal to normal (NN) intervals in ms
  + **SDNN:** Standard deviation of NN intervals in ms
  + **RMSSD:** RMS value of successive differences in ms
  + **PNN50:** Percentage of successive differences greater than 50 ms

*FreqDomainHrvMetrics*: this is frequency domain related output field containing following attributes:

* + **ULF:** Power in Ultra Low Frequency band in ms2
  + **VLF:** Power in Very Low Frequency band in ms2
  + **LF:** Power in Low Frequency band in ms2
  + **HF:** Power in High Frequency band in ms2
  + **LF/HF:** LF/HF ratio
  + **TOTPWR:** Total power in ms2

*percentCompleted*: Status of algorithm computation

*isHrvCalculated*: Flag that indicates if the content of the output is valid

HRV Algorithm’s status field is defined in “mxm\_hrv\_public.h” file distributed with Maxim Wellness Library:

**enum** MxmHrvRet {

*MXM\_HRV\_SUCCESS*,

*MXM\_HRV\_NULL\_PTR\_ERROR*,

*MXM\_HRV\_INVALID\_CONFIG\_ERROR*,

*MXM\_HRV\_NON\_POSITIVE\_SAMPLING\_PERIOD\_ERROR*,

*MXM\_HRV\_IBI\_PREP\_ERROR*,

*MXM\_HRV\_METRIC\_CALC\_ERROR*,

*MXM\_HRV\_ALREADY\_INITIALIZED\_ERR*,

*MXM\_HRV\_NOT\_INITIALIZED\_ERR*

}

## Respiration Rate Algorithm’s Output and Status Definitions

Respiration Rate Detection Algorithm’s output field is defined in “mxm\_respiration\_rate\_manager.h” file distributed with Maxim Wellness Library:

**struct** \_mxm\_respiration\_rate\_manager\_out\_data\_str {

**float** respiration\_rate; /\*\*< Output respiration rate with unit 'Breaths/Minute'\*/

**float** confidence\_level; /\*\*< Confidence level of the output in range of 0-100\*/

**boolean\_t** motion\_flag; /\*\*< True if motion detected\*/

**boolean\_t** ibi\_low\_quality\_flag; /\*\*< True if received IBI values are received with low quality\*/

**boolean\_t** ppg\_low\_quality\_flag; /\*\*< True if input PPG signal is noisy (DC Jumps, spikes) \*/

}

*respiration\_rate*: Output respiration rate with unit 'Breaths/Minute’.

*confidence\_level*: Confidence level of the output in range of 0-100.

*motion\_flag*: True if motion detected.

*ibi\_low\_quality\_flag*: True if received IBI values are received with low quality.

*ppg\_low\_quality\_flag*: True if input PPG signal is noisy (DC Jumps, spikes).

Respiration Rate Detection Algorithm’s status field is defined in “mxm\_respiration\_rate\_manager.h” file distributed with Maxim Wellness Library:

**enum** mxm\_repiration\_rate\_manager\_return\_code {

*MXM\_RESPIRATION\_RATE\_MANAGER\_SUCCESS*,

*MXM\_RESPIRATION\_RATE\_MANAGER\_INIT\_NULL\_PTR\_ERROR*,

*MXM\_RESPIRATION\_RATE\_MANAGER\_INIT\_ERROR*,

*MXM\_RESPIRATION\_RATE\_MANAGER\_RUN\_NULL\_PTR\_ERROR*,

*MXM\_RESPIRATION\_RATE\_MANAGER\_RUN\_ERROR*,

*MXM\_RESPIRATION\_RATE\_MANAGER\_END\_NULL\_PTR\_ERROR*,

*MXM\_RESPIRATION\_RATE\_MANAGER\_END\_ERROR*,

*MXM\_RESPIRATION\_RATE\_MANAGER\_GET\_VERSION\_NULL\_PTR\_ERROR*,

};

## Sleep Quality Assessment Algorithm’s Output and Status Definitions

Sleep Quality Assessment Algorithm’s output field is defined in “mxm\_sleep\_manager.h” file distributed with Maxim Wellness Library:

**struct** mxm\_sleep\_manager\_output\_data\_str {

mxm\_sleep\_manager\_sleep\_wake\_output\_status sleep\_wake\_decision\_status;

mxm\_sleep\_manager\_sleep\_wake\_output sleep\_wake\_decision;

**short** **int** sleep\_wake\_detection\_latency;

**float** sleep\_wake\_output\_conf\_level;

mxm\_sleep\_manager\_sleep\_phase\_output\_status sleep\_phase\_output\_status;

mxm\_sleep\_manager\_sleep\_phase\_output sleep\_phase\_output;

**float** sleep\_phase\_output\_conf\_level;

**float** hr;

**float** acc\_mag;

**float** interbeat\_interval;

};

**struct** mxm\_sleep\_manager\_output\_dataframe {

mxm\_sleep\_manager\_output\_data\_str \* output\_data\_arr;

**unsigned** **int** output\_data\_arr\_length;

uint64\_t date\_info

};

**Sleep/Wake Classification related outputs:**

*sleep\_wake\_decision\_status*: This enumeration is a flag indicating whether the sleep-wake decision is updated or not. Moreover, the sleep-wake-output status is set to NOT\_CALCULATED during the first N minutes (initialization period) defined by config *mxm\_sleep\_manager\_detection\_duration\_config*

*sleep\_wake\_decision*: Sleep/Wake state classifier output representing the status of sleep time instance:

MXM\_SLEEP\_MANAGER\_WAKE: short-term decision for wake state

MXM\_SLEEP\_MANAGER\_RESTLESS: short-term decision for restless state which corresponds to inconvenient portions during sleep rooting from small and non-periodic wiggles.

MXM\_SLEEP\_MANAGER\_SLEEP: short-term decision for sleep state

*sleep\_wake\_detection\_latency*:

*sleep\_wake\_output\_conf\_level*: Confidence level for the sleep/wake state classifier output.

**Sleep Phases Classification related outputs:**

*sleep\_phase\_output\_status*: Sleep Phases Output status as calculated or not yet calculated

*sleep\_phase\_output*: Sleep Phases Classification output, whose states can be:

MXM\_SLEEP\_MANAGER\_SP\_UNDEFINED: Sleep Phases Output is not ready

MXM\_SLEEP\_MANAGER\_SP\_WAKE: wake condition

MXM\_SLEEP\_MANAGER\_REM: Subject sleeps in in REM sleep

MXM\_SLEEP\_MANAGER\_LIGHT: Subject sleeps in in light sleep

MXM\_SLEEP\_MANAGER\_DEEP: Subject sleeps in in deep sleep

*sleep\_phase\_output\_conf\_level*: Sleep Phases Classification output confidence level.

*hr*: average heart rate value processed by SleepQA Algorithm

*acc\_mag*: average accelerometer magnitude value processed by SleepQA Algorithm

*interbeat\_interval*: average inter-beat interval value processed by SleepQA Algorithm

Sleep Quality Assessment Algorithm’s status field is defined in “mxm\_sleep\_manager.h” file distributed with Maxim Wellness Library:

**typedef** **enum** \_mxm\_sleep\_manager\_return {

*MXM\_SLEEP\_MANAGER\_SUCCESS*, // Success return code

*MXM\_SLEEP\_MANAGER\_INIT\_NULL\_PTR\_ERROR*, // NULL pointer error during initialization

*MXM\_SLEEP\_MANAGER\_INIT\_INVALID\_INPUT\_ERROR*, // Invalid Input error return code

*MXM\_SLEEP\_MANAGER\_ALGO\_INIT\_ERROR*, // Algorithm module initialization error return code

*MXM\_SLEEP\_MANAGER\_ALGO\_END\_ERROR*, // Algorithm module termination error return code

*MXM\_SLEEP\_MANAGER\_RUN\_NULL\_PTR\_ERROR*, // NULL pointer error during execution return code

*MXM\_SLEEP\_MANAGER\_RUN\_ALGO\_ERROR*, // Algorithm execution error return code

*MXM\_SLEEP\_MANAGER\_RUN\_ALGO\_AGE\_OR\_GENDER\_NOT\_DEFINED*, // Age or gender is not defined flag

}

## Stress Monitoring Algorithm’s Output and Status Definitions

Stress Monitoring Algorithm’s output field is defined in “mxm\_stress\_monitoring.h” file distributed with Maxim Wellness Library:

**struct** \_mxm\_stress\_monitoring\_run\_output {

boolean\_t stress\_class;

uint8\_t stress\_score;

float32\_t stress\_score\_prc;

}

*stress\_class*: Binary stress/non-stress output

stress\_score: Integer stress score output in interval [0 18]:

|  |  |
| --- | --- |
| **Score** | **Meaning** |
| 1-2 | Overwhelmed |
| 3-5 | Frustrated |
| 6-7 | Manageable Stress |
| 8-9 | Feeling Okay |
| 10-11 | Doing Great |
| 12-14 | Very Relaxed |
| 15-18 | Uncomfortably Relaxed |

Table : Stress Score Provision

*stress\_score\_prc*: Stress score output as percentage

Stress Monitoring Algorithm’s status field is defined in “mxm\_stress\_monitoring.h” file distributed with Maxim Wellness Library:

**enum** \_mxm\_stress\_monitoring\_return\_code {

*MXM\_STRESS\_MONITORING\_SUCCESS* = 0, //SUCCESS code

*MXM\_STRESS\_MONITORING\_NULL\_PTR\_ERR*, // NULL pointer to configuration structure

*MXM\_STRESS\_MONITORING\_INVALID\_CONFIG\_ERR*, // Configuration error

*MXM\_STRESS\_MONITORING\_RUN\_ERR*, // Algorithm process error

}

# Wellness Library API Function Definitions

## Algorithm Related Process

Maxim Wellness Library API functions are defined in “AlgoWrapper.h” file distributed with Maxim Wellness Library:

**mxm\_algosuite\_manager\_init( )** : Initialization call of the Maxim Wellness Library.

**Input:** init\_str : pointer to initialization structure of Maxim Wellness Library, described as [mxm\_algosuite\_init\_data](#_Wellness_Library_Configuration)

**Output:** status: pointer to Maxim Wellness Library return codes, described as [mxm\_algosuite\_return\_code](#_Wellness_Library_I/O)

**mxm\_algosuite\_manager\_run( )** : Method to dispatch processing of Maxim Wellness Library algorithms. Must be called with arrival of every new input sample. Outputs an output sample and a return code per input sample fed to method.

**Input:** data\_in\_str: pointer to input data sample of Maxim Wellness Library, described as

[mxm\_algosuite\_input\_data](#_Wellness_Library_I/O)

**Output:** data\_out\_str: pointer to Maxim Wellness Library output sample structure, described as [mxm\_algosuite\_output\_data](#_HRV_Algorithm’s_Output)

**Output:** status: pointer to Maxim Wellness Library status output sample structure, described as [mxm\_algosuite\_return\_code](#_HRV_Algorithm’s_Output)

**mxm\_algosuite\_manager\_end( ) :** End call of the Maxim Wellness Library.

**Input:** tobeDisabledAlgorithms: 8BIT filed to select algorithms to be disabled. Bit fields reserved for algorithms are explained in configuration definitions section.

**Output:** status: pointer to Maxim Wellness Library status sample structure, described as [mxm\_algosuite\_return\_code](#_HRV_Algorithm’s_Output).

***mxm\_algosuite\_manager\_get\_versions ( )*  :** returns the version information of Wellness Library.

**Output:** version\_str: pointer to Maxim Wellness Library version structure, described as:

struct mxm\_algosuite\_version\_str {

char version\_string[20]; //version in XX.XX.XX format

unsigned short int version; // The first number in XX.XX.XX format

unsigned short int sub\_version; //The second number in XX.XX.XX format

unsigned short int sub\_sub\_version; // The third number in XX.XX.XX format

}

***mxm\_algosuite\_calculate\_SQI ( )*** : Calculates the Sleep Quality Index related to subject's sleep. This function should be called after the detection of subject's wake-up for post processing of Sleep Quality assessment algorithm’s outputs.

**Input:** deep\_time\_in\_sec : time spent in DEEP state (in seconds) during the subject's sleep period

**Input:** rem\_time\_in\_sec : time spent in REM state (in seconds) during the subject's sleep period

**Input:** in\_sleep\_wake\_time\_in\_sec: total sleep time, time between Bedtime and Wake-Up time, duration in secs

**Input:** number\_of\_wake\_in\_sleep : count of wakes during the sleep period

**Output:** output\_sleep\_quality\_index: calculated Sleep Quality Index

## Authentication Related Process

Authentication to the SensorHub is a requirement for the Maxim Wellness Library. Authentication must take place before initialization of Algorithms within Wellness Library. Authentication Steps, which are also explained in integration document are:

1. Ask Authentication initials to the SensorHub, which is defined in **Max32664C/B API document**. Authentication initials is 12-character long byte array.
2. Call **mxm\_algosuite\_manager\_getauthinitials()** Wellness Library API function with the Authentication initials stream obtained in step1 from Sensorhub. Function returns out\_auth\_initials of the Wellness Library which is a byte stream of 12 characters.
3. Send the authentication array of Library obtained in step-2 to SensorHub via command defined in **Max32664C/B API document.**
4. Get the Two authentication keys of SensorHub via commands defined in **Max32664C/B API document.** Those Keys are:

* 12byte long auth\_str1[]
* 32byte long auth\_str2[]

1. Call **mxm\_algosuite\_manager\_authenticate()** Wellness Library API function with the Authentication Keys requested from SensorHub in step-4; auth\_str1[] and auth\_str2[] to complete authentication process, After completing Authentication, Algorithms within Wellness Library can be initialized and run.